## Slide 1
Experimental comparison of kinematics and control interfaces for laparoscope positioners Benoît Herman1,2, Alba Olías López1,3, Catherine Rasse4, Benoît Raucent1 1 Université catholique de Louvain, Center for Research in Mechatronics
2 Fonds de la Recherche Scientifique – FNRS
3 Universidad Politécnica de Madrid, Escuela Técnica Superior de Ingenieros Industriales
4 Université catholique de Louvain, Institute for Multidisciplinary Research in Quantitative Modelling and Analysis
NCTAM2012 – RMA, Brussels − 10.05.2012

## Slide 2: Imagine…
You drive a car
  On a sinuous road
  Backwards
  Looking only in a side mirror
  With you passenger moving the mirror now and then
  With one hand only on the steering wheel
  With a non constant steering ratio
  The other hand holding a cup full of boiling hot coffee
  And with one eye shut
Guess what?
    You’re performing laparoscopic surgery!
2

## Slide 3: Laparoscopic surgery
Minimally invasive technique
  Abdominal cavity explored with a rigid endoscope
  Proven benefits for patients: e.g. size of scars, traumatism and pain
  More challenging for surgeons: e.g. complex gestures, fatigue, stress
Manual camera handling by assistant
  Lack of coordination, wrong motions
  Image instability, tremors
  “Third hand” required
Passive instrument holders: not suitable
    Need for robotic assistance
3

## Slide 4: Existing robots
AESOP (Computer Motion)
EndoAssist (Armstrong Healthcare)
FreeHand (Prosurgics)
ViKY (Endocontrol)
LapMan (Medsys)
Vesalius (KULeuven)
4

## Slide 5: Performance assessment
No standard to measure performances…
Short clinical studies
  Proof of concept: most robots were introduced by this means
    Image stability, less lens-organs contacts, less fatigue, assistant free…
  Metric: operation duration
    Not reproducible
    Includes set up and break down
Ex vivo experiments [Yavuz 2001, Nebot 2003]
  More standardized motions
  Compares whole devices (AESOP vs. EndoAssist)
    What is the influence of each factor?
5

## Slide 6: Goal of this study
To assess the influence of two factors on motion duration (and intuitiveness)
  Kinematics (coordinate frame in which laparoscope motions are controlled)
  Number of controllable directions
Through a phantom experiment
  Task designed to avoid any bias
Using a single robotic platform to exclude all uncontrollable sources of variability between artifacts
  Maximum laparoscope velocity
  Acceleration and deceleration duration
  Reaction time of the interface
  Etc.
6

## Slide 7: Kinematics (1)
A laparoscope has 4 DOF
  Translation along longitudinal axis (zoom in and out)
  Self-rotation
  Rotation around two axes passing through the incision (left-right and up-down: pan motions of the image in the monitor)
Natural way to move it
  Intrinsic coordinate system: longitudinal axis + principal image directions
  Motion regulated by visual feedback on the monitor when performed by hand
AESOP does this
  Very intuitive image motions, as expected by the surgeon
  Need for 3 actuated joints for pan motions
  Kinematic coupling between actuators
7

## Slide 8: Kinematics (2)
To reduce complexity
  Hybrid coordinate system with one fixed and one mobile axis of revolution
  Decoupling possibility: one motor = one motion
    2 DOF require 2 motors (not 3)
2 main families
8
ViKY, EndoAssist, FreeHand, da Vinci S and Si
EVOLAP, LapMan, Vesalius*

## Slide 9: Control interface
“Natural” approach: real-time navigation
  Footpad
  Voice
  Head/eye motion
  Joystick on instrument
“Robotics” approach: target indication
  Instrument tracking (external or in image)
  Line of sight tracking
  Vocal control of a “cursor” on monitor
Characteristics
  Reaction time, order recognition rate, intuitiveness, comfort, …
    Depend mostly on implementation
  Number of controllable directions: usually Cartesian directions only!
    Additional diagonals, or omnidirectionality
9

## Slide 10: Methodology – Setup overview
monitor with red dot at center
pelvitrainer
joystick
laparoscope + camera
computer with dSPACE real-time controller + I/O’s
EVOLAP robot
laparoscope rotation device
10

## Slide 11: Methodology – EVOLAP overview
Local 1DOF zoom device: translation
Lockable articulated arm with passive RCM
Main 2DOF remote manipulator: angular motions
Decoupled 3DOF architecture
Instrument-mounted joystick omnidirectional and proportional
11

## Slide 12: Methodology – EVOLAP overview
Decoupled 3DOF architecture
 Static balancing with spring
 Back-drivability
Circular translations of the grasping point around the incision
No robot and arm motion during zoom
Intra-abdominal workspace ~ half-sphere
12

## Slide 13: Methodology – EVOLAP upgrade
Need for a third revolute DOF
  To move the laparoscope in all 3 coordinate frames
  Local self-rotation device replaces the unused zoom device
Control diagram update
  2 additional kinematics
  3 sets of directions
Use of Wii Nunchuck
  Good ergonomics
  Shorter learning curve than the instrument-mounted joystick
trocar
gears
elastic coupling
brushless DC motor + gearbox + encoder
13

## Slide 14: Methodology – Tasks and metrics
Point to point trajectory
Metrics
  Time for task completion
  Automatic stopwatch: starts with motion and stops when task completed
  Other real-time recordings for further investigation
14

## Slide 15: Methodology – Protocol
5 subjects
  PhD students in mechatronics
  No previous experience in robotic surgery
  Some are (very) good at video games
Randomization
  Kinematics in random order
  For each kinematics, joystick mode in random order
4 valid repetitions for each factors combination
  If learning effect, repeat until 4 repetitions are stable
15

## Slide 16: Results – Individual subjects
16

## Slide 17: Results – All subjects
17
2-way ANOVA
  Significant for both factors (p < .0001)
  No interaction between factors (p = 0.43)
Tukey HSD test for multiple comparisons
  ViKY slower than EVOLAP (p < .0005) and AESOP (p < .0001)
  EVOLAP a bit slower than AESOP but not significant (p = 0.1)
  Cartesian mode slower than with diagonals (p < .0001)
  Omnidirectional mode faster than with diagonals (p = 0.0025)

## Slide 18: Results – Influence of kinematics
18
ViKY
AESOP
EVOLAP

## Slide 19: Results – Influence of directions
19
Omni
Cartesian
+ diagonals

## Slide 20: Conclusions
New methodology for experimental assessment of motion performance of laparoscope positioners
Kinematics and number of controllable directions have significant influence on motion duration (and intuitiveness)
AESOP kinematics is optimal but EVOLAP offers a good balance between performance and complexity
Cartesian directions are not sufficient
Adding diagonals can be done easily with most modalities and improve performance
20

## Slide 21: What’s next?
Results should be confirmed with surgeons (including usability assessment)
Influence of angular speed should be assessed
  Possible plateau: too high speed might reduce accuracy
  Constant angular velocity does not lead to a constant image velocity!
Influence of kinematics on complex surgical tasks
  Self-rotation of the laparoscope changes the image orientation with respect to organs and instruments
  Could increase the difficulty to perform challenging gestures
Reach the market…
21

## Slide 22
Experimental comparison of kinematics and control interfaces for laparoscope positioners Benoît Herman1,2, Alba Olías López1,3, Catherine Rasse4, Benoît Raucent1 1 Université catholique de Louvain, Center for Research in Mechatronics
2 Fonds de la Recherche Scientifique – FNRS
3 Universidad Politécnica de Madrid, Escuela Técnica Superior de Ingenieros Industriales
4 Université catholique de Louvain, Institute for Multidisciplinary Research in Quantitative Modelling and Analysis
NCTAM2012 – RMA, Brussels − 10.05.2012